Update on liquid target production and isolation of [68Ga]GaCl3

T. Erikssona, J. Frigella, K. Gagnon1,a

aGE Healthcare, GEMS PET Systems, Uppsala, Sweden

Introduction

Attributed both to the ease of acquiring 68Ga by 68Ge/68Ga generators, and the approval and development of new tracers with theranostic relevance, there has been an ever-increasing use of 68Ga over the last decade. However, since 68Ge/68Ga generators are limited due to: 68Ga activity, the # of elutions, 68Ge decay, and the need for monitoring of 68Ge breakthrough, the direct cyclotron 68Zn(p,n)68Ga production route is an attractive alternative to ensuring a routine supply of 68Ga.

Two strategies exist for producing 68Ga via the 68Z(p,n)68Ga reaction – namely, liquid [1-4] and solid targets [5-6]. Liquid targets offer increased simplicity as they present a similar workflow to production of [18F]F-. If extraordinarily high 68Ga activity is desired, solid targets offer order magnitude increases in yield, albeit solid targets are typically met with a trade-off of increased requirements on infrastructure and/or local site expertise.

Be it liquid or solid, the direct cyclotron route requires an efficient means for purifying the 68Ga from the irradiated 68Zn and for controlling the proton energy to limit co-production of 67Ga to acceptable levels. Given the current cost of enriched 68Zn, recycling is considered unnecessary.

In this work, we provide an update on liquid target production of 68Ga, focusing on:

- Liquid target optimization/new target

- Chemical isolation of [68Ga]GaCl3

- Quality control, with a focus on radionuclidic purity and residual metals

Material and Methods

Irradiations were performed on a GE PETtrace cyclotron. The target media was comprised of zinc nitrate in dilute nitric acid, first using natZn, and ultimately moving to 68Zn. 68Ga yields were assessed by means of curve fitting/stripping of multiple dose calibrator measurements assayed for ~1-5 hrs (natZn) or ~1-3 hrs (68Zn).

The zinc and nitric acid concentrations were optimized considering 68Ga yield, 68Zn consumption, and ease of solution transfer to a hotcell.

Several liquid target designs were also considered and tested, including development of a new helium-cooling-free target with integrated energy degrader (Figure 1). Other than the PEEK fill lines, all target designs had only niobium as a contact material with the target media.

Figure 1. GE PETtrace 2mL 68Ga target with integrated energy degrader

Chemical isolation as [68Ga]GaCl3 was implemented on the GE FASTlab Developer Platform. In this scheme [7] (Figure 2), 68Ga is first trapped on hydroxamate resin (2 mL [~700 mg] ZR resin, Triskem), washed with 14 mL 0.1 M HNO3 and eluted as [68Ga]GaCl3 in 5 mL ~1.5 M HCl. The eluate is then efficiently trapped on TK200 resin (2 mL [~700 mg]) without the need for first increasing the HCl concentration. Finally, after purging with nitrogen, the trapped [68Ga]GaCl3 is eluted in 3 mL water.

Figure 2. [68Ga]GaCl3 isolation scheme

Quality control (QC) tests were performed on the extracted [68Ga]GaCl3 (e.g. half-life, radiochemical purity (RCP), etc), however, considerable effort was directed towards both the radionuclidic purity (RNP), and the residual metal content pre/post chemical purification.

The RNP was assessed for two different lots of 68Zn (Table 1), each lot of which was irradiated at 4 different nominal proton energies from ~13 to 15 MeV. The contribution of 66Ga and 67Ga was determined by gamma spectroscopy with a LaBr detector (efficiency calibrated using 137Cs and 152Eu sources), with samples analysed ~20-40 hrs post end-of-bombardment (EOB). No other non-6xGa gammas were identified.

Table 1. Experimental lots of enriched 68Zn

|  | 64Zn [%] | 66Zn [%] | 67Zn [%] | 68Zn [%] | 70Zn [%] |
| --- | --- | --- | --- | --- | --- |
| A | 0.76 | 0.61 | 0.27 | 98.33 | 0.03 |
| B | 0.03 | 0.16 | 0.62 | 99.16 | 0.03 |
| Nat | 49.17 | 27.73 | 4.04 | 18.45 | 0.61 |

In the context of metal content, the European Pharmacopoeia monograph for generator-based [68Ga]GaCl3 notes limits of 10 μg/GBq for each of iron and zinc. While it is possible to perform this test with adequate sensitivity using commercially available colorimetric test strips, for completeness, we have screened for ~70 elements via ICP-MS on a selection of sample retains of both irradiated (but not chemically processed) and also samples of chemically isolated [68Ga]GaCl3. Such samples have spanned the development phase. Use of the m/z=68 channel and appropriate isotopic abundance was used for analysis of zinc.

Results and Discussion

To optimize the production of 68Ga, it is not only a matter of optimizing the yield of 68Ga; one must also consider 68Zn consumption and ease of handling. As noted in Figure 3, (i.e. 15.2 MeV, 10 μA, tb = 30 minutes, natZn(NO3)2 in 0.2 M HNO3), the yield of 68Ga increased linearly with zinc concentration. However, such a result was limited only to low beam currents. As beam current was increased, pressures were undesirably elevated for the higher zinc concentrations. Furthermore, at high zinc concentrations, transfer to the hotcell was slow, if not impossible, for long transfer distances. We have therefore opted for a target media of 1.0 M [68Zn]Zn(NO3)2 in 0.2 M excess HNO3.

Figure 3. 68Ga yield vs. Zn concentration

Several technical challenges in existing chemistry methods for separating 68Ga from bulk 68Zn include, but are not limited to, high HCl concentration/volume, use of organic solvents, base-mediated pH adjustment, non-commercially available resins, etc. In addition, our goal was to implement the chemistry onto a FASTlab cassette, including not only the chemistry, but also column conditioning, and furthermore, leaving adequate space on the cassette to perform peptide labelling on the same cassette.

Table 2 presents the combined results of 68Ga production with the optimized target media in the new PETtrace 68Ga liquid target followed by [68Ga]GaCl3 separation on a FASTlab using the scheme presented above. We note that with a 1 hour irradiation using a liquid target, and ~23 minutes chemical processing, it is possible to achieve [68Ga]GaCl3 of similar chemical form with a 68Ga radioactivity in excess of a typical generator elution.

Table 2. Triplicate productions of [68Ga]GaCl3 [7]

| Parameter/Test | Value |
| --- | --- |
| Irradiation conditions | 14 MeV/60 min/30 μA |
| Total 68Ga at EOB | 4.4(1) GBq (120(4) mCi) |
| [68Ga]GaCl3 @ 30 min EOB | 2.5(1) GBq (68(4) mCi) |
| Radiochemical yield | 77(2) % |
| Half-life | 67.6(2) minutes |
| pH | <2 |
| Iron/zinc | Each <10 µg/GBq |
| RCP | 99(1) % |
| Appearance | Clear, colourless |

In the context of RNP, both 66Ga and 67Ga are co-produced. Table 3 reports the experimental 66Ga and 67Ga measurements for two commercial lots of 68Zn, irradiated at 4 different energies. As the irradiation time was not always the same, the saturation yields of 66Ga and 67Ga were calculated, and, for purpose of comparison, a 68-minute irradiation was assumed. These results confirm that the contribution of 66Ga is largely due to the contribution of 66Zn in the starting material, whereas the 67Ga content is energy dependent, with the contribution of 68Zn(p,2n)67Ga outweighing the contribution from 67Zn(p,n)67Ga for energies greater than ~13 MeV.

Table 3. RNP (assuming a 68 minute irradiation)

| E [MeV] | N | Lot 68Zn | 68Ga [%] | 67Ga [%] | 66Ga [%] |
| --- | --- | --- | --- | --- | --- |
| 13.1 | 3 | A | 99.94 | 0.0073(4) | 0.057(5) |
| 14.0 | 7 | A | 99.86 | 0.079(8) | 0.062(5) |
| 14.5 | 2 | A | 99.84 | 0.103(5) | 0.055(7) |
| 15.2 | 4 | A | 99.72 | 0.207(12) | 0.072(8) |
| 13.1 | 3 | B | 99.98 | 0.0130(3) | 0.0096(6) |
| 14.3 | 16 | B | 99.90 | 0.091(9) | 0.0099(8) |
| 14.5 | 3 | B | 99.89 | 0.104(6) | 0.0083(13) |
| 15.2 | 4 | B | 99.80 | 0.194(8) | 0.009(4) |

In assessing metal impurities, data is tabulated (Table 4) for both Zn and Fe (in μg/mL), as well as 8 other impurities (in ng/mL) for irradiation only “irrad“, and chemically isolated [68Ga]GaCl3 "chem”. These additional impurities are of interest due to known target contact materials, noted chemical variation between 68Zn lots, potential concern for competing for DOTA chelation, etc., however, we note that these are generally reduced to the single-to-low-tens of ng/mL following chemical purification. It is also important to emphasize that ICP-MS data presented below spanned the duration of the development period, some experiments of which were performed with ACS grade reagents, with such data included to emphasize robustness of the isolation scheme. Current levels of 68Zn are now routinely below 1 μg/mL, with further reduction demonstrated with additional column rinsing.

Table 4. Select elemental impurities [8]

|  |  | Zn Lot A |  | Zn Lot B |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Irrad (n=3) | Chem (n=9) | Irrad (n=5) | Chem (n=6) |
| μg | Zn | 48583(365) | 2.8(34) | 55183(1066) | 0.73(0.48) |
| /mL | Fe | 1.13(12) | 0.24(12) | 0.98(13) | 0.18(12) |
| ng | Ni | 1318(118) | 18(11) | 214(35) | 19(13) |
| /mL | Cu | 1615(121) | 8.8(61) | 1147(152) | 4.7(12) |
|  | Co | 85(74) | 0.28(19) | 3.8(11) | 0.150(55) |
|  | Cr | 37(20) | 22(24) | 46.0(87) | 10.2(52) |
|  | Y | 0.60(10) | 0.38(14) | 0.380(45) | 0.283(75) |
|  | Nb | 21(35) | 0.046(14) | 18(35) | 0.14(11) |
|  | Sn | 9(12) | 1.19(61) | 335(192) | 21(29) |
|  | Cd | 4.6(35) | 0.19(15) | 1940(87) | 0.22(12) |

Conclusion

We have presented a new PETtrace 68Ga target with an integrated energy degrader and we have developed and implemented a 2-column approach to isolate cyclotron-based [68Ga]GaCl3 on the FASTlab Developer platform. Results demonstrate it is possible to achieve [68Ga]GaCl3 with a 68Ga activity in excess of a typical generator elution. We note that the content of 66Ga and 67Ga may be reduced to acceptable levels [9] by selection of enriched 68Zn low in 66Zn, and by reduction of the proton energy. Analysis of residual metals supports the current applicability of cyclotron-based production of 68Ga and anticipated robustness as novel ligands and chelators are developed and introduced.

References

1.	A.T.J. Vogg, R. Lang, P. Meier-Böke, W. Scheel, S.N. Reske, B. Neumaier: 6th International Conference on Nuclear and Radiochemistry, Aachen, Germany, 2004.

2.	M. Jensen, J. Clark, 13th Workshop on Targetry and Target Chemistry, Roskilde, Denmark, 2010.

3.	E. Oehlke, C. Hoehr, X. Hou, V. Hanemaayer, S. Zeisler, M.J. Adam, T.J. Ruth, A. Celler, K. Buckley, F. Benard, P. Schaffer: Nucl. Med. Biol. 42, pp. 842-849, 2015.

4.	M.K. Pandey, N.R. Schmit, J.M. McEarchern, T.R. DeGrado, 16th Workshop on Targetry and Target Chemistry, Santa Fe, NM, USA, 2016.

5.	J.W. Engle, V.Lopez-Rodriguez, R.E. Gaspar-Carcamo, H.F. Valdovinos, M. Valle-Gonzalez, F. Trejo-Ballado, G.W. Severin, T.E. Barnhart, R.J. Nickles, M.A. Avila-Rodriguez: Appl. Radiat. Isot. 70, pp. 1792-1796, 2012.

6.	M. Lin, G. Waligorski, C.G. Lepera: Appl. Radiat. Isot. 133, pp. 1-3,  2018.

7.	M. Nair, S. Happel, T. Eriksson, M.K. Pandey, T.R. DeGrado, K. Gagnon: Eur. J. Nucl. Med. Mol. Imaging. 44 (Suppl 2), pp., S119-S956, 2017.

8. 	J. Frigell, K. Gagnon, 19th European Symposium on Radiopharmacy and Radiopharmaceuticals, Groningen, 2018.

9.	S.A. Graves, J. W. Engle, T.E. Eriksson, K. Gagnon, Annual meeting of the SNMMI, Philadelphia, 2018.