Indigenous development of Semi-Automated Solid target irradiation assembly compatible with PETtrace-800

Kanchan Kushwahaa, b,1, Pravind Malethaa, Sunil Kamblec, Jyothish Babuc, Tejindar Singh Sabharwalc, Kavindra Pathakc, Manish Kumarc, Nand Kishore Prasadc, Tushar Kanti Sahac, Sharmila Banerjee a, b

aRadiation Medicine Centre, BARC, TMH Annexe, Parel, Mumbai 400012, India

bHomi Bhabha National Institute, BARC, Anushakti Nagar, Mumbai 400094, India

cTechnical Physics Division, BARC, Trombay, Mumbai 400085, India

Introduction

64Cu has been known to be an effective radioisotope particularly in PET imaging studies.  A solid target assembly for the production of 64Cu via 64Ni (p, n)64Cu has been developed in-house in order to extend the utilization of the existing PETtrace-800 facility.

Materials and Method

Considering the compatibility requirements, the design was developed (fig-1). 16.5 MeV proton beam is allowed to sequentially incident on Havar foil, attenuation foil and 64Ni layer electroplated on base plate. Helium gas and water were used for heat dissipation (fig-2, table-1). The assembly body is made of aluminium and copper. Removal of base plate is carried out using semi-automation procedure. Stepper motor moves the base plate in horizontal direction and drops it in lead pot through a pipe, which will be carried out manually for chemical processing.

Fig-1 Schematic of solid target assembly

|  | Material | Working | Dimension (Diameter, thickness in mm) | Proton energy Attenuation (MeV), heat load (W)at 25 μA |
| --- | --- | --- | --- | --- |
| Havar | (Co, Cr, Ni, W, Mo, Mn, C, Fe) | To prevent Vacuum | 27, 0.025 | (16.5-16), 12.5 |
| Attenuation Foil | Al | To attenuate proton energy | 20, 0.5 | (16-12), 100 |
| Target layer | 64Ni | Target material | 8, 0.120 | (12-9), 75 |
| Base plate | Gold plated Cu | To support 64Ni layer | 25, 2 | (9-0), 225 |

Table-1 Dimension and Heat load of different part of solid target assembly

Fig-1 Solid target assembly

Fig-2 Schematic of solid target assembly

Results and Conclusions The use of Solid target assembly as described has been successfully demonstrated in the PETtrace-800 facility using proton beam of 2 µA for 30 m1. Further studies using higher beam current are under progress.

Reference

1)Production of 64Cu, using indigenously developed solid target assembly in Medical Cyclotron Facility in Radiation Medicine Centre (ESRR’18 – 19th Eur. Symp. Radiopharmacy Radiopharm. Apr 5-8,2018)